## ACR_Worksheet
ACR Clinical Test Checker – MRI Sequence Worksheet
Current Exam Type:
Max Slice Thickness (mm)
Max Gap (mm)
Max Pixel Area (mm²)
Exam Type | Body Region | Sequence Name / Type | Orientation | Slice Thickness (mm) | Gap (mm) | FOVp (mm) | FOVf (mm) | Phase Matrix (Np) | Freq Matrix (Nf) | Pixel (read, mm) | Pixel (phase, mm) | Pixel Area (mm²) | TR (ms) | TE (ms) | Flip Angle (°) | TI (ms) | # Averages | B value | Temporal Res (ms) | Pass / Review | Suggested Action
Note: Select the exam type in cell B3. The criteria above (B4–B6) come from the Criteria sheet and are based on ACR MRI accreditation examples. Adjust them only with physics/ACR review.

## Criteria
ACR-Style Criteria Table (Based on ACR MRI Accreditation Examples)
NOTE: Values are derived from ACR Head & Neck, Spine, MSK, Body, and MRA example parameter tables. Always confirm with the latest ACR documents and your local physicist.
Exam Type | Region / Notes | Max Slice Thickness (mm) | Max Gap (mm) | Max Pixel Area (mm²) | Comments
Brain | Routine brain / TIA | 5 | 2 | 1.2 | Structural sequences (FLAIR, T2, bright-fluid).
Brain DWI | Axial DWI + ADC | 5 | 2 | 4 | Brain DWI/ADC allows larger pixel area.
C-Spine | Cervical spine | 3 | 1 | 1 | High-detail sagittal/axial C-spine imaging.
L-Spine | Lumbar spine | 5 | 1.5 | 1.5 | Lumbar sagittal/axial sequences; ACR-style maxima.
Knee | MSK – knee | 4 | 1 | 0.6 | Internal derangement knee imaging.
Shoulder | MSK – shoulder | 4 | 0.8 | 0.8 | Shoulder MSK with fine spatial resolution.
Wrist/Hand | MSK – small joints | 3 | 0.5 | 0.3 | High-res wrist/small joint imaging.
Foot/Ankle | MSK – foot/ankle | 3 | 0.3 | 0.4 | Forefoot/ankle Morton-type high-res imaging.
Abdomen/Liver | Body – abdomen | 6 | 1.5 | 2 | Representative high-quality abdominal imaging limits.
Pelvis/Prostate | Body – pelvis | 5 | 1.5 | 1 | High-res pelvic/prostate sequences.
Carotid/Neck MRA | MRA – neck (2D TOF) | 2 | 0 | 1.1 | 2D TOF carotid/neck source images.